| Description | OQ Alarm: MX_580_MS_ConsFault1 |  |  |  |
| --- | --- | --- | --- | --- |
|  | Consecutive fault: Tube labeling, station 1 is faulty |  |  |  |
| Prerequisites | Machine is ready for operation \Operation\Monitoring, tube labeling\Settings\Monitoring, tube labeling A is on CHECK PATH |  |  |  |
| Step 1 | Manipulate the labeling area of 3 tubes, so the label cannot be applied, and mark them; then insert these tubes in sequence into track 1 Note value of \Operation\Monitoring, tube labeling\Settings\Monitoring, tube labeling A “Consecutive fault”: ____ Check Path Set “Consecutive fault” to 3 Check Path Transmit settings to the control system Press the START-button |  |  |  |
| Result 1 | Expected | Pass / Fail: Dev-No. | Date | Initials |
|  | Machine stops, after the third tube has passed the camera |  |  |  |
|  | The following message appears: |  |  |  |
|  | Consecutive fault: Tube labeling, station 1 is faulty |  |  |  |
| Step 2 | Press the START-button |  |  |  |
| Result 2 | Expected | Pass / Fail: Dev-No. | Date | Initials |
|  | The machine cannot be started. |  |  |  |
| Step 3 | Set “Consecutive fault” to the value noted in step 1 Transmit settings to the control system Press the RESET-button |  |  |  |
| Result 3 | Expected | Pass / Fail: Dev-No. | Date | Initials |
|  | Message is cleared |  |  |  |
|  | Machine status: ”Ready for operation” |  |  |  |